Multi-Drug 4 Drugs Rapid Test Panel

(Urine)

Package Insert

Instruction Sheet for testing of any combination of the following drugs:

COC/THC/AMP/OPI

A rapid test for the simultaneous, qualitative detection of multiple drugs and drug metabolites in human urine. For healthcare professionals including professionals at point of care sites.

For self-testing, not for diagnostic use.

【INTENDED USE】

The Multi-Drug Rapid Test Panel is a rapid chromatographic immunoassay for the qualitative detection of multiple drugs and drug metabolites in human urine at the following cut-off concentrations:

| Test | Calibrator | Cut-off (ng/mL) |
| --- | --- | --- |
| Cocaine (COC) | Benzoylecgonine | 300 |
| Marijuana (THC) | 11-nor-Δ9-THC-9 COOH | 50 |
| Methamphetamine (AMP) | d-Methamphetamine | 1,000 |
| Opiate (OPI) | Morphine | 2,000 |

Configurations of the Multi-Drug Rapid Test Panel come with any combination of the above listed drug analytes. This assay provides only a preliminary test result. A more specific alternate chemical method must be used in order to obtain a confirmed analytical result. Gas Chromatography/Mass Spectrometry (GC/MS) is the preferred confirmatory method. Clinical consideration and professional judgment should be applied to any drug of abuse test result, particularly when preliminary positive results are indicated.

【PRECAUTIONS】

- For healthcare professionals including professionals at point of care sites.

- For self-testing, not for diagnostic use. The test should remain in the sealed pouch until use.

- All specimens should be considered potentially hazardous and handled in the same manner as an infectious agent.

- The used test should be discarded according to local regulations.

【STORAGE AND STABILITY】

Store as packaged in the sealed pouch at 2-30°C. The test is stable through the expiration date printed on the sealed pouch. The Test must remain in the sealed pouch until use. DO NOT FREEZE. Do not use beyond the expiration date.

【SPECIMEN COLLECTION AND PREPARATION】

Urine Assay

The urine specimen should be collected in a clean and dry container. Urine collected at any time of the day may be used. Urine specimens exhibiting visible precipitates should be centrifuged, filtered, or allowed to settle to obtain a clear specimen for testing.

Specimen Storage

Urine specimens may be stored at 2-8°C for up to 48 hours prior to testing. For prolonged storage, specimens may be frozen and stored below -20°C. Frozen specimens should be thawed and mixed well before testing.

【MATERIALS】

Materials Provided

| - Test Panels | - Package Insert |
| --- | --- |

Materials Required But Not Provided

| - Timer | - Specimen collection containers |
| --- | --- |

【DIRECTIONS FOR USE】

Allow the test, urine specimen and/or controls to reach room temperature (15-30ºC) prior to testing.

- Bring the pouch to room temperature before opening it. Remove the test panel from the sealed pouch and use it within one hour.

- Remove the cap.

- With the arrow pointing toward the urine specimen, immerse the test panel vertically in the urine specimen for at least 10 to 15 seconds. Immerse the dipstick to at least the level of the wavy lines, but not above the arrow on the test panel.

- Replace the cap and place the test panel on a non-absorbent flat surface.

- Start the timer and wait for the colored line(s) to appear.

- The drug strip result should be read at 5 minutes. Do not interpret the result after 10 minutes.

【INTERPRETATION OF RESULTS】

(Please refer to the illustration above)

NEGATIVE:* A colored line appears in the control region (C) and another colored line appears in the test region (T). This negative result means that the concentrations in the urine sample are below the designated cut-off levels for a particular drug tested.

*NOTE: The shade of the colored lines(s) in the test region (T) may vary. The result should be considered negative whenever there is even a faint line.

POSITIVE: A colored line appears in the control region (C) and no line appears in the test region (T). The positive result means that the drug concentration in the urine sample is greater than the designated cut-off for a specific drug.

INVALID: No line appears in the control region (C). Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure. Read the directions again and repeat the test with a new test. If the result is still invalid, contact your manufacturer.

【QUALITY CONTROL】

A procedural control is included in the test. A line appearing in the control region (C) is considered an internal procedural control. It confirms sufficient specimen volume, adequate membrane wicking and correct procedural technique.

Control standards are not supplied with this kit. However, it is recommended that positive and negative controls be tested as good laboratory practice to confirm the test procedure and to verify proper test performance.

【LIMITATIONS】

- The Multi-Drug Rapid Test Panel provides only a qualitative, preliminary result. A secondary analytical method must be used to obtain a confirmed result. Gas Chromatography /Mass Spectrometry (GC/MS) is the preferred confirmatory method.1,2

- There is a possibility that technical or procedural errors, as well as interfering substances in the urine specimen may cause erroneous results.

- Adulterants, such as bleach and/or alum, in urine specimens may produce erroneous results regardless of the analytical method used. If adulteration is suspected, the test should be repeated with another urine specimen.

- A positive result does not indicate level or intoxication, administration route or concentration in urine.

- A negative result may not necessarily indicate drug-free urine. Negative results can be obtained when drug is present but below the cut-off level of the test.

- This test does not distinguish between drugs of abuse and certain medications.

- A positive test result may be obtained from certain foods or food supplements.

【PERFORMANCE CHARACTERISTICS】

Accuracy

% Agreement with GC/MS

|  | OPI 2,000 | COC 300 | AMP 1,000 | THC 50 |
| --- | --- | --- | --- | --- |
| Positive Agreement | 96.7% | 98.2% | 98.1% | 97.9% |
| Negative Agreement | 93.8% | 97.8% | 97.9% | 98.1% |
| Total Results | 95.2% | 98.0% | 98.0% | 98.0% |

% Agreement with Commercial Kit

|  | OPI 2,000 | COC 300 | AMP 1,000 | THC 50 |
| --- | --- | --- | --- | --- |
| Positive Agreement | * | >99.9% | >99.9% | >99.9% |
| Negative Agreement | * | >99.9% | >99.9% | >99.9% |
| Total Results | * | >99.9% | >99.9% | >99.9% |

*Note: Based on GC/MS data instead of Commercial Kit.

Precision

A study was conducted at three hospitals using three different lots of product to demonstrate the within run, between run and between operator precision. An identical card of coded specimens, containing drugs at concentrations of negative, 50% and 25% cut-off level, was labeled, blinded and tested at each site. The results gained ≧75% accuracy in ±25% cut-off level specimen and 100% accuracy in negative and ±50% cut-off level specimen.

Analytical Sensitivity

A drug-free urine pool was spiked with drugs at the listed concentrations. The results are summarized below.

| Drug Concentration Cut-off Range | OPI 2,000 |  | COC 300 |  | AMP 1,000 |  | THC 50 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | - | + | - | + | - | + | - | + |
| 0% Cut-off | 30 | 0 | 30 | 0 | 30 | 0 | 30 | 0 |
| -50% Cut-off | 30 | 0 | 30 | 0 | 30 | 0 | 30 | 0 |
| -25% Cut-off | 27 | 3 | 26 | 4 | 26 | 4 | 26 | 4 |
| Cut-off | 15 | 15 | 13 | 17 | 15 | 15 | 14 | 16 |
| +25% Cut-off | 5 | 25 | 3 | 27 | 3 | 27 | 3 | 27 |
| +50% Cut-off | 0 | 30 | 0 | 30 | 0 | 30 | 0 | 30 |
| +300% Cut-off | 0 | 30 | 0 | 30 | 0 | 30 | 0 | 30 |

Analytical Specificity

The following table lists the concentrations of compounds (ng/mL) that are detected as positive in urine by the Multi-Drug Rapid Test at 5 minutes.

| Analytes | conc. (ng/mL) | Analytes | conc. (ng/mL) |
| --- | --- | --- | --- |
| MORPHINE/OPIATE (OPI 2,000) |  |  |  |
| Codeine | 2,000 | Morphine | 2,000 |
| Ethylmorphine | 3,000 | Norcodeine | 25,000 |
| Hydrocodone | 50,000 | Normorphone | 50,000 |
| Hydromorphone | 15,000 | Oxycodone | 25,000 |
| Levorphanol | 25,000 | Oxymorphone | 25,000 |
| 6-Monoacetylmorphine | 3,000 | Procaine | 50,000 |
| Morphine 3--D-glucuronide | 2,000 | Thebaine | 25,000 |
| COCAINE (COC 300) |  |  |  |
| Benzoylecgonine | 300 | Cocaethylene | 20,000 |
| Cocaine HCl | 200 | Ecgonine | 30,000 |
| METHAMPHETAMINE (AMP 1, 000) |  |  |  |
| D,L-Amphetamine sulfate | 300 | Phentermine | 1,000 |
| L-Amphetamine | 25,000 | Maprotiline | 50,000 |
| (±) 3,4-Methylenedioxy amphetamine | 500 | Methoxyphenamine | 6,000 |
|  |  | D-Amphetamine | 1,000 |
| MARIJUANA (THC 50) |  |  |  |
| Cannabinol | 35,000 | △8-THC | 17,000 |
| 11-nor-△8-THC-9 COOH | 30 | △9-THC | 17,000 |
| 11-nor-△9-THC-9 COOH | 50 |  |  |

Effect of Urinary Specific Gravity

Fifteen (15) urine samples of normal, high, and low specific gravity ranges (1.005-1.045) were spiked with drugs at 50% below and 50% above cut-off levels respectively. The Multi-Drug Rapid Test was tested in duplicate using fifteen drug-free urine and spiked urine samples. The results demonstrate that varying ranges of urinary specific gravity do not affect the test results.

Effect of Urinary pH

The pH of an aliquoted negative urine pool was adjusted to a pH range of 5 to 9 in 1 pH unit increments and spiked with drugs at 50% below and 50% above cut-off levels. The spiked, pH-adjusted urine was tested with the Multi-Drug Rapid Test. The results demonstrate that varying ranges of pH do not interfere with the performance of the test.

Cross-Reactivity

A study was conducted to determine the cross-reactivity of the test with compounds in either drug-free urine or drug positive urine containing above related calibrator substances. The following compounds show no cross-reactivity when tested with the Multi-Drug Rapid Test at a concentration of 100 µg/mL.

Non Cross-Reacting Compounds

| Acetophenetidin | Cortisone | Zomepirac | Quinidine |
| --- | --- | --- | --- |
| N-Acetylprocainamide | Creatinine | Ketoprofen | Quinine |
| Acetylsalicylic acid | Deoxycorticosterone | Labetalol | Salicylic acid |
| Aminopyrine | Dextromethorphan | Loperamide | Serotonin |
| Amoxicillin | Diclofenac | Meprobamate | Sulfamethazine |
| Ampicillin | Diflunisal | Isoxsuprine | Sulindac |
| l-Ascorbic acid | Digoxin | d,l-Propanolol | Tetracycline |
| Apomorphine | Diphenhydramine | Nalidixic acid | Tetrahydrocortisone, |
| Aspartame | Ethyl-p-aminobenzoate | Naproxen | 3-acetate |
| Atropine | -Estradiol | Niacinamide | Tetrahydrocortisone |
| Benzilic acid | Estrone-3-sulfate | Nifedipine | Tetrahydrozoline |
| Benzoic acid | Erythromycin | Norethindrone | Thiamine |
| Bilirubin | Fenoprofen | Noscapine | Thioridazine |
| d,l-Brompheniramine | Furosemide | d,l-Octopamine | d,l-Tyrosine |
| Cannabidiol | Gentisic acid | Oxalic acid | Tolbutamide |
| Chloral hydrate | Hemoglobin | Oxolinic acid | Triamterene |
| Chloramphenicol | Hydralazine | Oxymetazoline | Trifluoperazine |
| Chlorothiazide | Hydrochlorothiazide | Penicillin-G | Trimethoprim |
| d,l-Chlorpheniramine | Hydrocortisone | Perphenazine | d,l-Tryptophan |
| Chlorpromazine | o-Hydroxyhippuric acid | Phenelzine | Uric acid |
| Cholesterol | 3-Hydroxytyramine | Prednisone | Verapamil |
| Clonidine | d,l-Isoproterenol |  |  |

【BIBLIOGRAPHY】

      - Hawks RL, CN Chiang. Urine Testing for Drugs of Abuse. National Institute for Drug Abuse (NIDA),Research Monograph 73, 1986.

      - Baselt RC. Disposition of Toxic Drugs and Chemicals in Man. 6th Ed. Biomedical Publ., Foster City,CA 2002.

Number:

Revision date: 2023-06-27